## Events
Event Number | 10X.Date_Time_Submitted | 104.Classification | 102.Location_Sub | 105.Narrative | 106.Tx_Technique | 107.Local_ID | Tag/Topic | 108.Reporter_Name | 202.Discoverer_Role | 205.Narrative_Supplemental | 201.Event_Title
78632 | 2023-09-05 |  | TKC-AR, The Kirklin Clinic at Acton Road | New start questionnaire said to schedule pt on TB. During the chart check I realized that pt was supposed to be scheduled on the IX
78630 | 2023-09-05 |  | HSROC, Hazelrig-Salter Radiation Oncology Center | At the time of physics second check, it was discovered the density overrides were not done correctly for a HN patient that had numerous metal clips and streaking artifacts. Specifically, the streaking artifacts surrounding the metal clips were not contoured/overridden to tissue and the metal was overridden to Ti when the HU was consistent with SS. The structure set was duplicated, the density overrides were fixed, and the plan was recalculated. Overall target coverage did not change appreciably, however, the dose was significantly hotter surrounding the clips. The physician elected to not modify the existing plan, but requested to have the plan with the correct density overrides prepared for treatment for proper record keeping.
69455 | 3/1/2023 10:35:26 AM | Near-miss | TKC-AR, The Kirklin Clinic at Acton Road | The patient has a contrast allergy listed in Powerchart and the sim was ordered with contrast. No mention of pre-medication for the sim. Message sent to MD to ask if patient was told to pre-medicate or if the order needs to be changed to without contrast. | Other |  | #sim | Stevie Sanders
70186 | 3/15/2023 10:44:52 AM | Operational/Process Improvement | HSROC, Hazelrig-Salter Radiation Oncology Center | Fiducial tracking y2 jaw set to +2.5 instead of -2.5. Therapist noticed at the end of fx 2 and changed parameters. | SRS/SBRT |  | #imaging | Chrystie Myers | Radiation Therapist |  | Therapy: Jaw tracking blades
70257 | 3/15/2023 7:33:32 PM | Therapeutic Radiation Incident | HSROC, Hazelrig-Salter Radiation Oncology Center | pt has been treated 4 times on the ix. Plan was actually calculated for stx. Therapist noticed when moding up plan that it has been overridden daily when calculated for stx. Therapist notified physics who had us halt treatment today while they recalculate. | 3D |  | #override | Chrystie Myers |  |  | Therapy/Physics: Machien Override
69611 | 3/2/2023 6:54:39 PM | Other Safety Incident | HSROC, Hazelrig-Salter Radiation Oncology Center | Order for patient was for a right upper extremity. Simulation therapists carefully removed sling and brace from patient and placed the arm akimbo per the standard setup for an extremity simulation. Attending physician came into the simulation room and stated brace and sling should not have been removed and the setup and procedure had to be repeated. No special instructions were provided to the simulation team on the order nor had the physician spoken with the simulation team concerning this patient. However, after the fact, the doctor told the simulation team he had obtained special permission to leave on the sling and brace from physics. Therefore, physician was aware of special instructions that were not relayed. | Other |  | #sim | Brandon Conant |  |  | Physicians: Sim Order
70684 | 3/23/2023 2:49:47 PM | Operational/Process Improvement | HSROC, Hazelrig-Salter Radiation Oncology Center | Proton patient was listed under Hazelrig as the location (instead of proton) it was imported and sent to the physician from the photon side. dosi happen to notice it when they was doing a random search in whiteboard. The patient had been turned over two days previously for planning, | Particles (Protons) |  | #whiteboard | A. Paige Dalton | Dosimetrist |  | Dos: Whiteboard
70740 | 3/24/2023 8:53:46 AM | Operational/Process Improvement | TKC-AR, The Kirklin Clinic at Acton Road | While checking a new start chart for an extended distance femur I realized the iso was mid femur and the orientation was head first. It was scanned and planned head first but should have been feet first. | 3D |  | #sim | Stevie Sanders | Radiation Therapist | Spoke with physician who said that he will put orders in for feet first when the TV is at the level of the femoral heads or lower. The dosimetry group has discussed when we need to rotate patient orientation and how to accomplish that. Information was added to a dosimetry document that cover common clearance issues. | Dos/Therapy: Orientation
70814 | 3/27/2023 9:15:11 AM | Other Safety Incident | HSROC, Hazelrig-Salter Radiation Oncology Center | Therapist not on that treatment machine printed the chart on Thursday. Therapist noticed on the will call ticket that the patient was had a cardiac stent and also noticed that there was bolus that needed to go under the mask. Therapy Assistant scheduled patient on Friday morning and did not notice that the patient had a cardiac stent/loop recorder as it was not in the patient encounter. Therapist on machine checked chart Mon morning when the patient was set to start at 6:30 and did not notice the cardiac stent physics report when checking the chart. A different therapist walked around and asked the more experienced therapist who had printed the chart what their idea was about placing bolus under the mask. After answering, it was asked if she knew that the pacemaker had an implanted device. The therapist who had checked the chart did not notice. They had to make the patient wait until the device could be interrogated. After waiting for 1.5 hours for the interrogators to arrive, they found out that it did not really need to be interrogated. The patient ended up leaving and not getting a treatment. | IMRT/VMAT |  | #pacemaker | Chrystie Myers | Physicist,Radiation Therapist | Patient was delayed about 30 minutes; did not actually go home as the report says. The ICD report had been written with generic language saying it should be interrogated prior to treatment. It is a loop recorder and does not need to be. | Physics: ICD
70921 | 3/29/2023 9:34:25 AM | Near-miss | TKC-AR, The Kirklin Clinic at Acton Road | The yellow sheet was labeled that patient would be breath hold, but the prescription had "none" listed under gating. Verified with DOS and they changed prescription to include breath hold. | 3D |  | #Gating |  |  |  | RX: Gating
69647 | 3/3/2023 3:41:29 PM | Other Safety Incident | HSROC, Hazelrig-Salter Radiation Oncology Center | A quick turn around TBI patient was approved, printed and scheduled. During physics review it was found that the spinal cord dose needs a minimum of 1080 but was currently only 1053cGy. Dosimetry was asked to replan which involved updating 6-7 plans as well as reprinting. This was also additional work for physics and decreased the time frame therapy had to review plan information. | IMRT/VMAT |  | #normaltissue | Korie Meyers | Physicist | Spinal cord was underdosed according to typical protocol. Physics was unable to get a response from dosimetry in an appropriate timeframe and contacted the physician. She wanted a replan based on that. Dosimetry later said they were previously instructed by the physician that 85% coverage (slight underdosing) was acceptable, but this was not documented anywhere. | Dos/Physician : Doses
71048 | 3/30/2023 2:42:34 PM | Therapeutic Radiation Incident | UAB Proton Center | During the patients imaging the robotic couch would not move. MCR came into the treatment room and re set some computers - during this time the patient became upset and it was decided to do her treatment under anesthesia. After the patient fell asleep additional images were acquired, the patient began to wake up again requiring more anesthesia and more imaging. The last image set reflects a vertical shift that did not align with the patients treatment fields correctly. Previously therapists have been very mindful of the patients chin, which was aligned correctly. The thorax misalignment was brought to the treating therapists attention by the DOD reviewing images. Therapists reviewed the images and time stamps and agreed that the anatomy was not aligned. The POD was notified. Therapists agreed to be more vigilant during image review prior to treatment and to modify filters to reflect the anatomy more clearly. | Particles (Protons) |  | #override | AS |  |  | Therap: Imaging
71079 | 3/31/2023 7:20:06 AM | Operational/Process Improvement | HSROC, Hazelrig-Salter Radiation Oncology Center | Patient scheduled for STx machine on 3/28 to start on Monday, 4/3/23 per the Scheduling task dropped by dosimetry. Therapist was checking new start patients for Monday, 4/3 on 3/31 and noticed that this patient's plan was calculated for Ethos. We called the dosimetrist who said yes they changed the plan/machine but forgot to tell the therapists. Called patient to give a different start time for treatment. | IMRT/VMAT |  | #machineswitch | Nathan Jordan |  |  | Dos: Machine Reassignment
69826 | 3/8/2023 9:17:18 AM | Operational/Process Improvement | HSROC, Hazelrig-Salter Radiation Oncology Center | Patient was to start a boost. The plan was not changed in ARIA to indicate a plan revision was starting. Plan was ready and printed and placed in chart weeks prior. Therapy never checked the chart. Therapist was filling in on the machine and noticed when checking images that the patient had a plan revision and was able to check prior to starting. It would have been optimal to know in ARIA that there was a plan revision that needed to be checked. Also, the patient was adaptive and it had not been coordinated with the attending physician to attend the new adaptive plan session. Luckily, another physician stepped in to cover. | IMRT/VMAT |  | #replan | Chrystie Myers |  |  | Therapy: plan revision
69828 | 3/8/2023 9:19:30 AM | Other Safety Incident | HSROC, Hazelrig-Salter Radiation Oncology Center | RX called for osld day 1. They were not done. Treating therapist noticed on day 2 and physics did the readings then. | IMRT/VMAT |  | #ivd | Chrystie Myers |  |  | All: Process for OSLD
71637 | 4/11/2023 1:17:26 PM | Operational/Process Improvement | UAB Proton Center | Revision plan not treatment approved morning of start. | Particles (Protons) |  | #urgent | Ali Schnoor | Radiation Therapist
71643 | 4/11/2023 1:20:54 PM | Operational/Process Improvement | UAB Proton Center | Bilateral expander patient eRx ordered CBCT weekly. Expander patients should be CBCT daily, requested patient alert. | Particles (Protons) |  | #imaging | Ali Schnoor | Radiation Therapist |  | Physician: incorrect erx?
71645 | 4/11/2023 1:25:28 PM | Operational/Process Improvement | UAB Proton Center | Plan not treatment approved morning of new start appointment. | Particles (Protons) |  | #urgent | Deanna Vail | Radiation Therapist
71744 | 4/12/2023 4:48:46 PM | Therapeutic Radiation Incident | HSROC, Hazelrig-Salter Radiation Oncology Center | Physician stated during clinical setup for neck keloid "No Bolus". The therapist put 106 SSD on the electron setup sheet. During planning bolus was added and it was planned using the 106 SSD. The treat sheet, and RX did show bolus, but the patient was treated with no bolus. | Electrons |  | #electron | Dara Burns | Dosimetrist |  | Physicians/Therapy/Dosimetry: Bolus
71753 | 4/12/2023 6:27:12 PM | Operational/Process Improvement | UAB Proton Center | Patient is a photon to proton transfer. When checking the consents therapist noticed that "left breast/ chest wall" was selected. Patient is being treated for bilateral breast. | Particles (Protons) |  | #Rx | A Schnoor | Radiation Therapist |  | Physician/Therapy: Consent
71811 | 4/13/2023 12:57:47 PM | Therapeutic Radiation Incident | HSROC, Hazelrig-Salter Radiation Oncology Center | PT has matching CW/SCLAV/Electron fields for her LT CW tx. On this day, KVs were taken to align the CW Iso. Then shifts were made to SCLAV iso, the AP SCLAV PORT was filmed, couch parameters were captured and then treated. Today, on Thursday, I started tx with the CW iso, made shifts to SCLAV section and noticed the VRT was off by +3 cm. Upon further inspection, I realized that shifts made from CW to SCLAV on tues were made in the opposite direction for VERTICAL SHIFTS AND TREATED. Physics was notified immediately. | 3D |  | #override | Shashank Singh |  |  | Therapy: Inappropriate shifts
71937 | 4/17/2023 10:01:36 AM | Operational/Process Improvement | UAB Proton Center | Plan not treatment approved the morning the patient was scheduled to start. Called physics for them to approve plan. | Particles (Protons) |  | #urgent | Ali Schnoor |  |  | Physics: TX approval
72146 | 4/19/2023 1:12:19 PM | Operational/Process Improvement | UAB Proton Center | During physician review of SSU setup, it was mentioned there was a discussion between physicians to change the patient's setup to frogleg. Therapy was not notified, and the order was not updated. Device then needed to be remade. | Particles (Protons) |  | #sim | Faye Watson | Radiation Therapist |  | Physician: Sim order,adding to Sim Task Force. also presented at faculty meeting.
72160 | 4/19/2023 3:09:52 PM | Near-miss | HSROC, Hazelrig-Salter Radiation Oncology Center | Patient under treatment for breast on ix using feathering technique. This specific plan requires multiple electron blocks for each feathering fraction. 4/18 was scheduled to be the first feathering fraction however this fraction of treatment was delayed by one day due to RTT's not receiving the electron block template for the first shift. RTT's had to contact dosimetry in order to access required template. | Electrons,3D |  | #electron | Ashlyn Gallahar | Radiation Therapist |  | Dosimetry: Electron Cutouts
72199 | 4/20/2023 11:41:31 AM | Therapeutic Radiation Incident | TKC-AR, The Kirklin Clinic at Acton Road | 21 MU delivered without gating. | IMRT/VMAT |  | #Gating | Stevie Sanders |  |  | Therapy: Gating on IX
72290 | 4/24/2023 7:24:11 AM | Near-miss | HSROC, Hazelrig-Salter Radiation Oncology Center | Patient was simulated 4/19 and scheduled to start treatment on STx on 4/24 by sim therapists. STx was indicated on sim setup notes and in Processmaker. Patient was planned for iX instead, and schedule new start task was dropped, indicating patient would start 4/24 on iX but not mentioning patient's pre-existing appointment on Stx. On 4/21, patient was scheduled to start treatment on iX. On 4/24, therapist checking chart noticed patient was scheduled on both iX and STx for that day. | IMRT/VMAT |  | #machineswitch | Madelyn Wine | Radiation Therapist |  | Dosimetry: Machine Scheduling
72309 | 4/24/2023 9:24:52 AM | Operational/Process Improvement | HSROC, Hazelrig-Salter Radiation Oncology Center | Need to start patient same day: Therapists received a Vocera at 8am on 4/24/23 from the planning Dosimetrist to start a patient in the afternoon on 4/24/23. The patient was simulated at 9am on 4/20/23. Plan was ready for treatment when the message was sent. | 3D |  | #urgent | Nathan Jordan |  |  | Physicians: Short timeline
72311 | 4/24/2023 9:33:06 AM | Operational/Process Improvement | HSROC, Hazelrig-Salter Radiation Oncology Center | Short notice in scheduling: Schedule patient next business day because chemo starts. Therapists were contacted on Friday, 4/21/23 at 1524 to schedule patient on Monday, 4/24/23. The patient was simulated on 4/13/23. | IMRT/VMAT |  | #urgent | Nathan Jordan |  |  | Physician: Short Timeline,Short timeline
72400 | 4/25/2023 9:15:50 AM | Operational/Process Improvement | HSROC, Hazelrig-Salter Radiation Oncology Center | Short notice for scheduling: patient was simulated on 4/6/23. Message sent at 3:22pm on 4/24/23 to start ASAP. Patient scheduled to start treatment on 4/25/23 at 10am. Plan was not treatment approved by 8:15am as of 4/25/23. Had to coordinate CIED interrogation for pacemaker. | SRS/SBRT |  | #urgent | Nathan Jordan |  |  | Physicians: Short Timeline
72720 | 4/28/2023 4:12:11 PM | Therapeutic Radiation Incident | HSROC, Hazelrig-Salter Radiation Oncology Center | Patient transitioned from protons to photons due to insurance delays. Physician added PTV contours but did not apply any expansion for CTV_Low. Error was not caught until patient had initiated treatment, initial plan with no PTV margin treated for 13 fractions. | 3D,IMRT/VMAT |  | #PTV | Joe Harms |  |  | Physicians: Contours,Physician: contours
71203 | 4/3/2023 1:55:46 PM | Operational/Process Improvement | HSROC, Hazelrig-Salter Radiation Oncology Center | Therapist was doing chart check on patient's chart. It was noticed that the PSQA report in powerchart contained an LDR eye plaque document versus the PSQA. Physics was notified and correct documentation was uploaded. | Total body irradiation (TBI) |  | #documentation | Shanna Wilson |  |  | Phys: Document Upload
71210 | 4/3/2023 3:29:35 PM | Therapeutic Radiation Incident | HSROC, Hazelrig-Salter Radiation Oncology Center | Hypofractionated brain (6 Gy x 5 fractions) treated on Edge. Fraction 5 (last fraction) treated on April 3. Several hours after fraction 5, the image reviewer noted offset of between CBCT and planning CT. Attending radiation oncologist investigated and notified physics. The CBCT shift given in Offline Review was Vrt: -0.05 cm, Lng: -0.93 cm, Lat: -0.74 cm, Rtn: -0.1 deg, Pitch: 0.3 deg, Roll: -0.1 deg. The SGRT logs were recovered and fraction 5 log was compared with the previous fractions. Prior to capture of the reference surface, the average offset from the DICOM reference surface for the first 4 fractions was Vrt: -0.12 cm, Lng: 0.05 cm, and Lat: -0.04 cm. On 04/03/2023, the offset prior to capture of the reference surface was Vrt: -0.14, Lng: 1.03 cm, and Lat: 0.63 cm. The SGRT logs confirmed that there was an unintended shift in the lateral and longitudinal direction. The hypothesis is that the image registration was accidentally moved after image alignment and review when the mouse pointer was in the CBCT coronal view prior to applying the couch shifts. Dose reconstruction in the TPS indicates that this incident is reportable to the state and radiation safety was notified, | IMRT/VMAT,SRS/SBRT |  | #override | Richard Popple |  |  | Treatment: Shifts applied incorrectlly
71275 | 4/4/2023 4:32:46 PM | Operational/Process Improvement | HSROC, Hazelrig-Salter Radiation Oncology Center | New start task had dropped for therapist/assistant to contact patient with new start date and time. However, therapist was unable to reach the patient. When reviewing the chart for additional contact numbers, therapist noticed the prescription in the chart did not match the site and patient currently being scheduled. Therapist accessed the !Plan documents and opened the delete folder to determine if the folder for the patient was sent over incorrectly or if chart was assembled incorrectly. It was determined that the iPlan document contained the wrong RX. The RX was for one patient and the yellow page and dosimetry treatment plan was for patient being scheduled. | IMRT/VMAT |  | #documentation | Shanna Wilson | Radiation Therapist | previous patient's RX was accidentally added to documentation. Therapy caught this in pretreatment chart check. | Dos/Phys: Plan documentation incorrect
71400 | 4/6/2023 11:48:37 AM | Operational/Process Improvement | TKC-AR, The Kirklin Clinic at Acton Road | MD was planning on treating patient on the Ethos but the sim order did not include that. The patient said something about her treatments had to be downtown. We asked the MD and that is when they told us the plan was to treat on the Ethos. | Other |  | #Rx | Stevie Sanders | Radiation Therapist |  | Physician: Orders
71492 | 4/7/2023 10:52:49 AM | Operational/Process Improvement | UAB Proton Center | eRx states (x2) no prior RT. Patient has prior prostate RT from 2011 in Aria. | Particles (Protons) |  | #prior | Deanna Vail | Radiation Therapist
71507 | 4/7/2023 2:02:17 PM | Operational/Process Improvement | HSROC, Hazelrig-Salter Radiation Oncology Center | On 4/5 pt finished Ethos tx on cspine/thorax. Therapists threw away head mold and told pt to take facemask home. Hours after pt finished the dosimetrist called to say that the pt needed to return possibly Friday to start Lattice tx; however, dosimetrist was trying to get in touch with attending to confirm. The attending was out of town the entire week of 4/7 and he must be present for tx. Therapists had to call pt and tell him to keep his facemask and bring it back for his treatment when we rescheduled it for 4/12 (when attending would be present). Therapists then took head mold out of trash and thoroughly sanitized it. Therapists then found out on 4/7 that the Lattice plan is 6 ARCs which means it will take approximately 45min-1 hour to calculate. The patient is on 6 liters of oxygen with a high flow oxygen mask during his prior treatments. | IMRT/VMAT |  | #boost | Laurel Robinson | Radiation Therapist |  | Physician/Dosimetry: Scheduling
71449 | 4/7/2023 8:15:07 AM | Operational/Process Improvement | HSROC, Hazelrig-Salter Radiation Oncology Center | During chart check for patient, the therapist noticed that the plan check (Mobius) for 4.1LHipInf was not attached to the powerchart documentation. Physics was notified and it was attached. | IMRT/VMAT |  | #documentation | Shanna Wilson | Radiation Therapist
73256 | 5/10/2023 9:15:58 AM | Near-miss | HSROC, Hazelrig-Salter Radiation Oncology Center | Patient was scheduled to start a boost 5/11, patient had re-sim for boost on 5/8. Therapist who was doing chart check noticed there was no gating/breath hold selected in rx however, therapist knew the patient had been treated with breath hold throughout primary course of treatment. Vocera was sent to MD, resident and dosimetrist to confirm accuracy of rx. Dosimetrist and MD verified that the rx was incorrect. MD stated "I think she was planned breath hold. Template must not have had that. I did not catch." | 3D |  | #Rx | Ashlyn Gallahar |  |  | Physician: eRX
73414 | 5/12/2023 10:17:30 AM | Operational/Process Improvement | TKC-AR, The Kirklin Clinic at Acton Road | Sim order said not radiosurgery. Patient said it was 5 fx radiosurgery and would have to have his tx downtown. The MD was on vacation and does not have a resident. Clinic note said 5-6 Gy with 5 fx. | Not Applicable |  | #sim | Stevie |  |  | Physician: Sim Order
73620 | 5/17/2023 10:41:26 AM | Operational/Process Improvement | HSROC, Hazelrig-Salter Radiation Oncology Center | Treatment machine was unaware patient was starting a plan revision. Once we tried pulling up patient's treatment plan that's when we noticed their plan was still planning approved. Spoke with dosimetry about issue and a task was never dropped for physics to check chart. QA was done prior. | IMRT/VMAT |  | #replan | Jessie Lacy | Radiation Therapist |  | Dosimetry: Carepath
73627 | 5/17/2023 11:13:47 AM | Near-miss | HSROC, Hazelrig-Salter Radiation Oncology Center | Patient's prescription stated phase based gating, but the yellow sheet said 0-90. Spoke to dosimetry about confusion and it was confirmed to be 0-90. Reached out to doctor and they changed prescription to show free breathe. | 3D |  | #phase | Jessie Lacy |  |  | Physicians: eRX Gating
73687 | 5/18/2023 7:46:35 AM | Operational/Process Improvement | HSROC, Hazelrig-Salter Radiation Oncology Center | Patient sent a message in the portal inquiring about their start date. Plan was still not ready and therapist called to tell patient to expect a phone call in a week or so with a start date. Patient informed therapist in passing that was perfectly fine and they start chemo in the morning. Therapist then called the doctor and resident to confirm that it was okay for the patient to start chemo and not radiation. It was confirmed that both chemo and radiation should begin the same day. Resident and Doctor spoke with VA to get chemo pushed and therapist was told to place patient on the schedule for the same day that the VA scheduled the new chemo appointment for. | IMRT/VMAT |  | #chemocoord | Sydni McLaughlin
72807 | 5/2/2023 10:14:16 AM | Near-miss | HSROC, Hazelrig-Salter Radiation Oncology Center | Breast patient had triggered imaging in Rx but it was not needed during treatment. No Breath Hold listed in Rx and confirmed by physician prior to treatment. Patient was planned with a breath hold and confirmed by physics. In conclusion the Rx said triggered imaging and no breath hold but was treated with no triggered imaging and breath hold. We have asked physician to modify the RT prescription. | SRS/SBRT |  | #Rx | Alan Gigis | Physicist,Radiation Therapist
72804 | 5/2/2023 8:52:04 AM | Operational/Process Improvement | HSROC, Hazelrig-Salter Radiation Oncology Center | Therapist was reviewing information on patient prior to calling patient and scheduling them. RX on patient stated continuous, consecutive days versus once daily, consecutive days. Therapist informed DOS, POD and MD that RX needed to be changed and plan placed back into treatment approved status. | IMRT/VMAT |  | #Rx | Shanna Wilson
73853 | 5/22/2023 2:09:32 PM | Operational/Process Improvement | HSROC, Hazelrig-Salter Radiation Oncology Center | Patient had missed several treatments and needed an updated schedule. Therapist noticed that patient was to get a boost per original prescription and asked the dosimetrist if patient was still to receive one. Dosimetrist responded with yes and patient needed to be re-ct'd after 3600 cGy. This was conveyed on whiteboard, but therapy was not aware of this. | IMRT/VMAT |  | #replan | Sydni McLaughlin
73811 | 5/22/2023 7:54:51 AM | Operational/Process Improvement | HSROC, Hazelrig-Salter Radiation Oncology Center | While therapist was checking chart for new start, the therapist noticed CT-Sim was completed prior to physician's signature on informed consent document. | Not Applicable |  | #sim | Howell Kyle Oglesby
73819 | 5/22/2023 9:13:32 AM | Operational/Process Improvement | HSROC, Hazelrig-Salter Radiation Oncology Center | Patient was scheduled to start on the iX. The day the patient was to start, the therapist was checking to see if the plan was treatment approved and ready to check when they noticed in treatment prep that the patient's plan was scheduled for the Stx instead. | 3D |  | #machineswitch | Ashlyn Lambert//Sydni McLaughlin
73899 | 5/23/2023 10:14:09 AM | Operational/Process Improvement | HSROC, Hazelrig-Salter Radiation Oncology Center | The physicist that checked the plan did not add to/finish the glucose monitor physics consult and did not complete the special physics consult task in aria. The consult and plan approval had to be completed same day of treatment. | SRS/SBRT |  | #SPC | Viscariello
73895 | 5/23/2023 9:43:02 AM | Operational/Process Improvement | HSROC, Hazelrig-Salter Radiation Oncology Center | Triggered imaging task was not dropped for a prostate SBRT with triggered imaging in the Rx | SRS/SBRT |  | #SPC
74002 | 5/24/2023 5:02:13 PM | Operational/Process Improvement | HSROC, Hazelrig-Salter Radiation Oncology Center | Keloid electron new start is planned with 9e- energy and 1cm bolus. Original RT Prescription stated 6e- energy and 0.5cm bolus. We alerted the treating physician to correct the RT Prescription. The only item changed was the thickness of the bolus from 0.5cm to 1cm. As the therapists are performing their chart check they notice that the energy on the RT Prescription and plan do not match. The therapist alerts the physician again to have the RT Prescription revised. Then the therapist alerts the Physicist to complete Treatment Approval. This caused about a 50 minute delay from the patient's scheduled time. | Electrons |  | #Rx | Nathan Jordan
74100 | 5/26/2023 9:08:35 AM | Therapeutic Radiation Incident | HSROC, Hazelrig-Salter Radiation Oncology Center | I was treating patient with a lateral shift that was out of tolerance that I overrode. When treating the 4th field of her plan there was a couch rotation. When I overrode the lateral shift I selected all rather than only selecting the lateral shift that was out of tolerance and then treated that field so the patient was treated at a 0 degree rotation rather than 90 degrees. The next field also had a couch rotation. The other treating therapist stopped me from selecting all when I was overriding on this field and the appropriate couch rotation was made before treating this field. | 3D |  | #override | Howell Kyle Oglesby
72894 | 5/3/2023 10:04:39 AM | Other Safety Incident | UAB Proton Center | Gating not engaged in TPS for breath hold abdomen. Had to close treatment and edit. | Particles (Protons) |  | #Gating | Deanna Vail | Radiation Therapist | This item is listed on the dosimetry checklist; however, it is so rarely used on the proton side that it is easily overlooked. | Dos: Gating Checkbox
72896 | 5/3/2023 10:07:46 AM | Therapeutic Radiation Incident | UAB Proton Center | Patient's right breast was not taped superiorly as it was at simulation. Written details were missing from simulation setup sheet. Discussed with therapy group, all proton liver treatments will have the standard setup of taping RT breast superior, if applicable. | Particles (Protons) |  | #sim | James Kraus/Deanna Vail |  |  | Simulation: Setup
72969 | 5/4/2023 8:20:27 AM | Near-miss | HSROC, Hazelrig-Salter Radiation Oncology Center | Rx for an SBRT says continuous consecutive | SRS/SBRT |  | #Rx | Natalie Viscariello |  |  | Physicians: eRX
73040 | 5/5/2023 10:23:47 AM | Near-miss | HSROC, Hazelrig-Salter Radiation Oncology Center | For a bilateral breast treatment with RNI on the right side, the CTV_Chestwall_R extended outside of the CTV^Right^Total and PTV^Right (laterally 1-2cm near the axilla for 4-5 slices). Xcheck plan quality check flagged the CTV contour extending outside PTV, but dosimetrist did not follow up with physician. The initial treatment plan was optimized to cover CTV_Chestwall_R. During physics plan review, physicist reached out to physician to confirm they wanted this region covered even though it extended outside PTV. Physician stated CTV_Chestwall_R needed to be cropped back inside PTV^Right and plan had to be reoptimized | 3D,IMRT/VMAT,SRS/SBRT,Particles (Protons) |  | #PTV | Joe Harms |  |  | Physicians: Contours X-check
74917 | 6/14/2023 2:31:51 PM | Operational/Process Improvement | HSROC, Hazelrig-Salter Radiation Oncology Center | Patient was rect'd for boost on 5/30 and physician did not begin working on the boost plan until 6/5. The boost was supposed to start on 6/7. Patient had to go on break until 6/12 when the boost plan was ready. | IMRT/VMAT |  | #replan
| 6/14/2023 2:31:51 PM | Operational/Process Improvement | HSROC, Hazelrig-Salter Radiation Oncology Center | Patient was rect'd for boost on 5/30 and physician did not begin working on the boost plan until 6/5. The boost was supposed to start on 6/7. Patient had to go on break until 6/12 when the boost plan was ready. | IMRT/VMAT |  | #revision
74923 | 6/14/2023 3:54:56 PM | Operational/Process Improvement | HSROC, Hazelrig-Salter Radiation Oncology Center | The physician prefers to specify OLSD placement for keloids. At 01:25, therapist messaged physician (who was also DOD) that they were getting patient on table for keloid treatment and that the physics resident was present for OSLD placement. They asked the physician to come to IX, and the physician said "Coming in 10 min" at 01:25. The therapists deemed it more efficient to keep the patient set-up than to take down the patient/cone and try to treat another patient in a 10 minute slot. The physician walked into the IX vault at 01:59, 34 minutes later. The patient sat on the table the entire 34 minutes because all involved parties expected the physician to show up at any time. | Electrons |  | #delay
| 6/14/2023 3:54:56 PM | Operational/Process Improvement | HSROC, Hazelrig-Salter Radiation Oncology Center | The physician prefers to specify OLSD placement for keloids. At 01:25, therapist messaged physician (who was also DOD) that they were getting patient on table for keloid treatment and that the physics resident was present for OSLD placement. They asked the physician to come to IX, and the physician said "Coming in 10 min" at 01:25. The therapists deemed it more efficient to keep the patient set-up than to take down the patient/cone and try to treat another patient in a 10 minute slot. The physician walked into the IX vault at 01:59, 34 minutes later. The patient sat on the table the entire 34 minutes because all involved parties expected the physician to show up at any time. | Electrons |  | #electron #rush
74970 | 6/15/2023 2:16:31 PM | Operational/Process Improvement | HSROC, Hazelrig-Salter Radiation Oncology Center | Patient had a plan revision load on the machine. Therapists were not notified that a plan revision was starting. Dosimetry did not notify therapist or change the appointment to the green revision status on ARIA. Dosimetry checked that they notified therapists and changed appointment status on their checklist but did not actually do those things. | IMRT/VMAT |  | #replan
| 6/15/2023 2:16:31 PM | Operational/Process Improvement | HSROC, Hazelrig-Salter Radiation Oncology Center | Patient had a plan revision load on the machine. Therapists were not notified that a plan revision was starting. Dosimetry did not notify therapist or change the appointment to the green revision status on ARIA. Dosimetry checked that they notified therapists and changed appointment status on their checklist but did not actually do those things. | IMRT/VMAT |  | #scheduling
74411 | 6/2/2023 2:31:59 PM | Therapeutic Radiation Incident | HSROC, Hazelrig-Salter Radiation Oncology Center | The treatment plan has 18 needles and total 4 fraction. In the third fraction, we tried to connect the transfer tubes with the needles and found one needle short. After careful checking, we found the 7th needle was pulled out together with the starlet. When the physician in the OR, he only had 17 titanium needles available which can be locked with the template. The 7th needle was used but could not be locked on the template due to different size. We cancelled the treatment of this fraction and modified the plan without Needle #7 (it only had 4.2 seconds in the original plan). After careful review, the new plan with 2 fractions was used to treat the patient. NO harm was caused to the patient. | HDR |  | #brachy
75162 | 6/20/2023 10:30:30 AM | Near-miss | UAB Proton Center | Plan made it through physics review then physician wanted to change CTVs (made them larger) once the plan was in 'Physician plan review'. | Particles (Protons) |  | #PTV
| 6/20/2023 10:30:30 AM | Near-miss | UAB Proton Center | Plan made it through physics review then physician wanted to change CTVs (made them larger) once the plan was in 'Physician plan review'. | Particles (Protons) |  | #target
| 6/20/2023 10:43:11 AM | Near-miss | UAB Proton Center | No Diagnosis coded entered in Aria. Caught by dosimetry when plan is printed/approved. | Particles (Protons) |  | #Rx
75163 | 6/20/2023 10:43:11 AM | Near-miss | UAB Proton Center | No Diagnosis coded entered in Aria. Caught by dosimetry when plan is printed/approved. | Particles (Protons) |  | #Sim
| 6/26/2023 4:53:27 PM | Operational/Process Improvement | TKC-AR, The Kirklin Clinic at Acton Road | Patient was approved by physician, but the chart was not written up for two weeks. It was discovered when the planning dosimetrist was reviewing process maker tasks. The patient was will call so no cues for any other departments to follow up about a start date. It is easy to miss a patient chart write up because of the automatic peer review comments in PM. A peer review comment will override previous notes on that status. Additional amendments that would help are some safety nets for patients that haven't been started. | 3D |  | #whiteboard
| 6/26/2023 5:02:47 PM | Operational/Process Improvement | HSROC, Hazelrig-Salter Radiation Oncology Center | Dosimetrist started the DOs and pushed through to physics with the following notes "DO^Ti and DO^SS related to the pacemaker Please adjust". The DOs were approved, but additional questions were brought up prior to plan start. A second opinion was sought and the DOs were added and adjusted to provide better overrides near and in the target volume. | SRS/SBRT |  | #DO
| 6/26/2023 5:13:47 PM | Operational/Process Improvement | HSROC, Hazelrig-Salter Radiation Oncology Center | Dosimetrist started the DOs and pushed to physics with the following comments "DO^Ti_5104 created. Please add additional DO for muscle/adipose and inside the bone". They were updated and approved, but the plan was not able to calculate due to some outstanding pixels with no density. Dosimetrist tried to return the DOs back to physics for correction according to our protocol ( urgent turn around+two or more DO in the target volume), but physics wanted dosimetry to correct. This led to delay in planning time. | IMRT/VMAT |  | #DO
| 6/26/2023 5:20:54 PM | Operational/Process Improvement | HSROC, Hazelrig-Salter Radiation Oncology Center | Dosimetry was concerned about a patient's weight on the STX. We have been told a range is 300-330lbs depending on the patient's position on the table. The patient was about 300lbs and a brain case. When dosimetry reached out to the POD to double check that this would be okay, the physicist said that the weight limit was actually 500lbs. Dosimetry expressed concerns about that weight limit being incorrect. It wasn't understood that the weight limit was actually 300-330lbs until the dosimetrist involved another physicist. It could be helpful to have this information updated with the physics team. | IMRT/VMAT |  | #tabletop
| 6/28/2023 5:24:50 PM | Operational/Process Improvement | UAB Proton Center | Peds anesthesia not treatment approved morning of new start appointment. | Particles (Protons) |  | #urgent
| 6/28/2023 5:27:32 PM | Operational/Process Improvement | UAB Proton Center | Proton abdomen treatment. Slide bar location not indicated on simulation setup record. I had a discussion with simulation staff on the day discovered. | Particles (Protons) |  | #simrecord
| 6/28/2023 5:34:49 PM | Operational/Process Improvement | UAB Proton Center | CSI Align structure not included with data set. | Particles (Protons) |  | #structure
| 6/29/2023 10:56:31 AM | Operational/Process Improvement | HSROC, Hazelrig-Salter Radiation Oncology Center | Patient was sent to dosimetry with top priority for planning. However, a couple of the OARs were not contoured properly including the liver (which the TV was located in) and the bowel. The dosimetrist sent the patient back to the physician for contour corrects which causes delay in planning start. It could be helpful to have attending or residents to run x-check prior to sending to dosimetry. it would have at least caught the liver in this case. | IMRT/VMAT |  | #autoplanning
| 6/29/2023 11:14:28 AM | Operational/Process Improvement | HSROC, Hazelrig-Salter Radiation Oncology Center | A process maker status was not updated after a patient's simulation was complete. The comment section read "please import;phase selection." the patient was not imported because dosimetry is not notified by a comment change. This was only caught when a dosimetry team member was clearing off patients who were never actually simulated. | IMRT/VMAT |  | #whiteboard
| 6/29/2023 5:15:23 PM | Other Safety Incident | UAB Proton Center | Staging patient is allergic to contrast, but sim was ordered with contrast. The patient was not aware and not told to pre-medicate. | Not Applicable |  | #sim
| 6/30/2023 5:26:52 PM | Other Safety Incident | HSROC, Hazelrig-Salter Radiation Oncology Center | SBRT patient's shifts were grossly off on the yellow sheet. After filming the patient and seeing that shifts were wrong, physics looked up the patient's plan and confirmed they were indeed inaccurate and that the user origin from CT was not correct. MD verified alignment and we proceeded with treatment. | SRS/SBRT |  | #isocenter
74456 | 6/5/2023 8:36:31 AM | Operational/Process Improvement | HSROC, Hazelrig-Salter Radiation Oncology Center | MD prescription indicated cylinder size to be 35 cm instead of 3.5 cm for brachytherapy treatment. Prescription revised before treatment. | HDR |  | #Rx | Brandon Conant
74575 | 6/6/2023 4:09:16 PM | Operational/Process Improvement | UAB Proton Center | New start was cancelled due to collision issues and told to expect a call when the replan was complete. Therapy was not notified revision was complete and starting the next day. | Particles (Protons) |  | #replan | Lacey Funderburk
74576 | 6/6/2023 4:14:51 PM | Operational/Process Improvement | UAB Proton Center | Upon opening CSI plan, only 1.2 boost plan was available to treat, but 1.1 plan had not finished. Investigated with AM POD and determined PM POD and MD had decided to hold 1.1 plan and tx 1.2 boost. Therapy/AM POD/DOD were not notified of changes. | Particles (Protons) |  | #boost | Deanna Vail/Bobbie Paulat
74578 | 6/6/2023 4:16:30 PM | Operational/Process Improvement | TKC-AR, The Kirklin Clinic at Acton Road | Gating option not selected for BH left breast new start for TB. | 3D |  | #Gating | Deanna Vail
74689 | 6/8/2023 2:01:05 PM | Operational/Process Improvement | TKC-AR, The Kirklin Clinic at Acton Road | Patient consented for radiation therapy to thorax. Plan named 2.1 Abodmen. Reconsent needed on 1st day of treatment. | IMRT/VMAT |  | #sim | Deanna Vail
74689 | 6/8/2023 2:01:05 PM | Operational/Process Improvement | TKC-AR, The Kirklin Clinic at Acton Road | Patient consented for radiation therapy to thorax. Plan named 2.1 Abodmen. Reconsent needed on 1st day of treatment. | IMRT/VMAT |  | #Sim
74701 | 6/8/2023 4:06:51 PM | Operational/Process Improvement | HSROC, Hazelrig-Salter Radiation Oncology Center | Patient was at HSROC to get CT scan for proton therapy. Patient was simulated while sucking on a peppermint. Noticed buy dosimetrist during import, where an unknown object with 367 HU was noted in the lower oral cavity. Possibly a choking hazard, as the patient was in an mask at the time. Definitely suboptimal for proton planning. | Particles (Protons) |  | #Sim | James Kraus
74701 | 6/8/2023 4:06:51 PM | Operational/Process Improvement | HSROC, Hazelrig-Salter Radiation Oncology Center | Patient was at HSROC to get CT scan for proton therapy. Patient was simulated while sucking on a peppermint. Noticed buy dosimetrist during import, where an unknown object with 367 HU was noted in the lower oral cavity. Possibly a choking hazard, as the patient was in an mask at the time. Definitely suboptimal for proton planning. | Particles (Protons) |  | #Sim
74706 | 6/8/2023 6:11:05 PM | Unsafe condition | HSROC, Hazelrig-Salter Radiation Oncology Center | Density override procedure not being followed. Reviewed plans from two patients today from the same planner where the density overrides for the first plan were performed and approved by the dosimetrist and the second plan had multiple high density artifacts (dental fillings, bone cement) around and in the target and no density override was performed. | Not Applicable |  | #Density
74706 | 6/8/2023 6:11:05 PM | Unsafe condition | HSROC, Hazelrig-Salter Radiation Oncology Center | Density override procedure not being followed. Reviewed plans from two patients today from the same planner where the density overrides for the first plan were performed and approved by the dosimetrist and the second plan had multiple high density artifacts (dental fillings, bone cement) around and in the target and no density override was performed. | Not Applicable |  | #DO
| 7/10/2023 9:30:50 AM | Operational/Process Improvement | TKC-AR, The Kirklin Clinic at Acton Road | Sim was ordered for a 4d scan. When the doctor came in to review sim they realized we had only done 4d scan and asked us to add expiratory bh scan. |  |  | NA
| 7/13/2023 4:52:16 PM | Operational/Process Improvement | HSROC, Hazelrig-Salter Radiation Oncology Center | Patient was scheduled per the encounter instructions on the IX. However, during the chart check by the therapist, it was noted that the patient had been planned for the STX. Patient had to be contacted and rescheduled due to incorrect machine information. |  |  | #scheduling
| 7/14/2023 2:11:44 PM | Near-miss | TKC-AR, The Kirklin Clinic at Acton Road | The new start questionnaire said to schedule pt for TB @ AR. When we pulled up plan on TB, it was asking for a machine override. Pt was supposed to be scheduled on IX. Moved pt's schedule to IX. |  |  | #scheduling
| 7/19/2023 1:53:34 PM | Near-miss | UAB Proton Center | Physician marked contours as "ready for planning" and did not do careful review of contours. The following issues were noticed and sent back to the physician for edits: "There is a "spare" CTV_Low contour at z=10.4 that is not connected with anything else, on z= 5.8 the structure doesn't avoid the mandible when on the slices before and after it, it does. ON the CTV_High on slice z= -4.6 there is a "spare" ditzel just anterior to the airway that is unconnected to the rest of the volume. Please adjust and send back to dosimetry". This delays planning and patient start. |  |  | #target
| 7/19/2023 5:05:27 PM | Operational/Process Improvement | HSROC, Hazelrig-Salter Radiation Oncology Center | In power chart senior physicist did not sign Jr Physicist documentation concerning pacemaker and did not give high/med/low recommendation causing delay in processing and treatment |  |  | #pacemaker
| 7/2/2023 2:09:19 PM | Therapeutic Radiation Incident | HSROC, Hazelrig-Salter Radiation Oncology Center | Bowel was not properly segmented by autoseg. Resulting approved plan had exceedingly high dose to a portion of the bowel that was not drawn or reflected in DVH/X-Check. | IMRT/VMAT |  | #autoplanning
| 7/25/2023 4:57:12 PM | Unsafe condition | TKC-AR, The Kirklin Clinic at Acton Road | Physics task was on the board to do a 'Treatment plan review' for a breast plan. All documents were printed and it appeared ready for standard procedure. After completing the check; the plan was treatment approved. After the fact, the physicist was notified that a revision plan was ready. There was no communication to physics that a revision was being planned, therefore the original plan was checked for no reason. In this case, the patient was not scheduled for same-day treatment; however, this is often the case and this could've resulted in treating with the un-revised plan. This may be just a fluke as it is a breast VMAT plan on the truebeam, and dosimetry was unaware they needed to use partial arcs. |  |  | #planningtechnique #communication
| 7/25/2023 5:02:07 PM | Unsafe condition | TKC-AR, The Kirklin Clinic at Acton Road | Physics task was on the board to do a 'Treatment plan review' for a breast plan. All documents were printed and it appeared ready for standard procedure. After completing the check; the plan was treatment approved. After the fact, the physicist was notified that a revision plan was ready. There was no communication to physics that a revision was being planned, therefore the original plan was checked for no reason. In this case, the patient was not scheduled for same-day treatment; however, this is often the case and this could've resulted in treating with the un-revised plan. This may be just a fluke as it is a breast VMAT plan on the true beam, and dosimetry was unaware they needed to use partial arcs. Second case in same day; both cases required replans for clearance purposes. |  |  | #planningtechnique #communication
| 7/25/2023 6:37:34 PM | Operational/Process Improvement | TKC-AR, The Kirklin Clinic at Acton Road | Complex density override was requested, but there were targets drawn on both breathe hold and free breathe scans. It was not indicated which structure set to perform DO on. Also, it was just two surgical clips in breast tissue which is arguably not a complex case - this will be addressed in the upcoming revision of the density overrides procedure. |  |  | #DO
| 7/26/2023 2:57:35 PM | Unsafe condition | TKC-AR, The Kirklin Clinic at Acton Road | A breath hold breast planned for the True Beam was planned with 5 full Arcs. High risk of collision due to lateral shift of 7. |  |  | #planningtechnique
| 7/26/2023 2:59:06 PM | Unsafe condition | TKC-AR, The Kirklin Clinic at Acton Road | A breath hold breast planned for the True Beam was planned with 5 full Arcs. High risk of collision due to lateral shift of 8. |  |  | #planningtechnique
| 7/27/2023 11:31:37 AM | Operational/Process Improvement | TKC-AR, The Kirklin Clinic at Acton Road | Patient planned with rapid arc but the treatment technique on the prescription said tangents. |  |  | #Rx
| 7/28/2023 10:59:02 AM | Near-miss | TKC-AR, The Kirklin Clinic at Acton Road | Prostate patient was imaged to check bladder filling. Bladder was not full. Therapists planned to get pt down while bladder filled but changed their minds and let pt lay on table while we waited. Pt's chart was closed out of the 4d console. The next patient arrived for treatment and we checked pt in on the 4d console leaving the 2nd patient's name highlighted. When it was time to recheck the original patient, therapist clicked the treat button without changing which patient's name was highlighted. A CBCT was done on the 2nd patient's chart when we then realized that it was the wrong patient pulled up. The chart was then closed and the correct patient was pulled up and proceeded with treatment. |  |  | #wrongpatient
| 7/28/2023 2:48:30 PM | Operational/Process Improvement | UAB Proton Center | Dr forgot to order mouthpiece, and patient informed therapist after devices were already made. |  |  | #sim
76903 | 7/28/2023 4:53:05 PM | Near-miss | TKC-AR, The Kirklin Clinic at Acton Road | Pt had Brain sim with contrast scheduled. He has iodine listed as an allergy. While doing the will call sheet, pt said contrast made his bp drop before when he had contrast. Pt was not premedicated. Messaged doctor and they said not to give contrast. |  |  | #sim
| 7/3/2023 4:40:49 PM | Other Safety Incident | HSROC, Hazelrig-Salter Radiation Oncology Center | SPC task for implanted devices was not completed on physics list, powerchart was not signed by a physicist, but plan reviewed and treat approved - devices were noted in powerchart plan review note. Patient had neurostimulator and glucose pump. Neurostimulator is not removable, but the glucose pump is removable. Physicist believed both devices were removable. Plan was done with 15X, 3D, 100 cGy./1fx plan. Plan proceed with 15X due to low MU | 3D |  | #pacemaker
4151299 | 7/31/2023 12:22:57 PM | Operational/Process Improvement | TKC-AR, The Kirklin Clinic at Acton Road | Pelvis patient with no instruction on bladder filling. Instructions were for full bladder at time of sim but not for time of treatment. |  |  | #sim
| 7/31/2023 9:09:25 AM | Operational/Process Improvement | TKC-AR, The Kirklin Clinic at Acton Road | Prescription notes state - 64 yo M w/ cT2N1b SCC of the anus. Planning concurrent chemoRT w/ 5-FU/MMC for primary plan but then state - No prior RT. 39 yoF with cT3N1 anal SCC. Planning concurrent chemoradiation for sequential boost plans. The patient is in fact a 39 yo female. |  |  | #Rx
76926 | 7/31/2023 9:19:19 AM | Operational/Process Improvement | TKC-AR, The Kirklin Clinic at Acton Road | Sim order for a base of skull. Order was directed as a CSI setup with conventional devices and 3mm slices but needed to be radiosurgery skull with encompass mask, slightly lower scan parameters, and 1mm slices. |  |  | #sim
| 7/5/2023 7:26:02 AM | Operational/Process Improvement | HSROC, Hazelrig-Salter Radiation Oncology Center | TBI submitted for plan review the afternoon before first treatment, but plan was signed by physician the day before. Dose rate on the lung fields was not set correctly and the plan had be remeasured for QA. Generally tight timeline for TBI which take significant physics time to check and can be error-prone. | Total body irradiation (TBI) |  | #urgent
| 7/5/2023 9:02:40 AM | Near-miss | HSROC, Hazelrig-Salter Radiation Oncology Center | Rx and plan laterality did not match. (r vs l lung). Caught in physics check, quickly resolved by physician. |  |  | #rx
| 7/6/2023 3:38:24 PM | Other Safety Incident | HSROC, Hazelrig-Salter Radiation Oncology Center | Patient fell in treatment room due to foot "falling asleep" after adaptive treatment. |  |  | #fall
| 8/1/2023 12:13:23 PM | Unsafe condition | HSROC, Hazelrig-Salter Radiation Oncology Center | Received a vocera from dosi, therapist, and physician 10min before patient was set to be treated requesting a plan check be done. Perineum/groin boost for TSET case. No task or mention of this plan starting from dosi or MD. MD didn't approve plan until 5 min before treatment, at which point dosi started printing plan documents (there were no plan or cutout documents uploaded). Rejected plan check and advised delaying patient due to unsafe, rushed condition. Treatment delayed to follow day. Come to find out during setup verification of cutout 30 min after incident (just verification of cutout shape, no intention to treat today), the wrong cutout was poured. There's a reason we don't want to rush these things... |  |  | #urgent
77009 | 8/1/2023 12:13:23 PM | Unsafe condition | HSROC, Hazelrig-Salter Radiation Oncology Center | Received a vocera from dosi, therapist, and physician 10min before patient was set to be treated requesting a plan check be done. Perineum/groin boost for TSET case. No task or mention of this plan starting from dosi or MD. MD didn't approve plan until 5 min before treatment, at which point dosi started printing plan documents (there were no plan or cutout documents uploaded). Rejected plan check and advised delaying patient due to unsafe, rushed condition. Treatment delayed to follow day. Come to find out during setup verification of cutout 30 min after incident (just verification of cutout shape, no intention to treat today), the wrong cutout was poured. There's a reason we don't want to rush these things... |  |  | #urgent
77013 | 8/1/2023 12:19:27 PM | Unsafe condition | HSROC, Hazelrig-Salter Radiation Oncology Center | Dosi messaged yesterday saying there will be a revision plan starting today at 3pm. Dosi sent plan for review by physician around 5pm yesterday. I advised dosi that machine schedule is completely full the day the revision is set to start and it will be a problem doing QA prior to treatment. Dosi vocera'd MD with this info. Plan wasn't approved until 10am today and documents not printed until 11am. Now in a situation of a rushed plan check for physics and the therapists, which creates opportunities for mistakes to occur. In the future, if a similar situation is encountered, suggest the plan must be approved and printed the night prior to starting. |  |  | #revision
77466 | 8/10/2023 10:54:12 AM | Unsafe condition | HSROC, Hazelrig-Salter Radiation Oncology Center | Density override was requested for targets near the pelvis. Primary concern was that GTV were drawn outside the body. I'm not familiar with current dosimetry workflow, is dosimetry supposed to check for this before or after? A secondary (and less severe) concern is that wire is being overridden outside the body. This is unecessary unless the wire crosses into the patients anatomy; DO procedure is being revised and it will be explicitly stated that objects 'outside the body contour' that will not be present at treatment should not be overridden. |  |  | #target
77477 | 8/10/2023 11:52:00 AM | Operational/Process Improvement | HSROC, Hazelrig-Salter Radiation Oncology Center | SBRT Left Lung New Start on Edge. The shifts from sim on yellow sheet were 15.5 cm posterior, 23.9 cm superior, and 9.0 cm left. The actual shifts were 5.0 cm posterior, 2.5 cm superior, and 9.0 cm left. |  |  | #isocenter
77519 | 8/10/2023 6:33:19 PM | Operational/Process Improvement | HSROC, Hazelrig-Salter Radiation Oncology Center | I "found" an upper body alpha cradle in the closet that belonged to the patient that has currently been under treatment for 11 fractions already. Looking back at the encounters, the pt had only one CT scan done with this "found" cradle, dated 6/20/23. For these 11 fxs, a different cradle has been used but that cradle actually belongs to this patient. The CT scan that was used for planning involved the "found" cradle. Somehow, somewhere, another cradle was made for the pt on 5/31/23 but there is no documentation of it anywhere in ARIA. Images have been approved daily for every fraction but the wrong cradle is being used to set up the patient. The cradles do look similar but they are not the same. We are going to use the "found" cradle from tomorrow onwards (8/11/23). |  |  | #immobilization
77579 | 8/11/2023 5:41:04 PM | Operational/Process Improvement | TKC-AR, The Kirklin Clinic at Acton Road | Breast simulation order did not state Ethos protocol. Patient made therapists aware that the physician told her she would have to be treated downtown to ensure treatment was on a particular machine. Therapy messaged physician to confirm this. |  |  | #sim
77580 | 8/11/2023 5:45:15 PM | Operational/Process Improvement | TKC-AR, The Kirklin Clinic at Acton Road | Simulation for rectal cancer. During zoom coverage of the scan (that was already completed), physician asked therapists if gross disease was wired. It had not been since it was not ordered. |  |  | #sim
77616 | 8/14/2023 12:45:34 PM | Operational/Process Improvement | HSROC, Hazelrig-Salter Radiation Oncology Center | A CNS patient was simmed and added in Processmaker as will call with surgery coordination. The patient was then scheduled for treatment on the Edge before the plan was completed. There was no communication with dosimetry (vocera from MD, scheduling email sent) about the date and time of treatment, or the priority level of the patient. The MD spoke with dosimetry the day before scheduled treatment inquiring about the plan, causing a rush in treatment planning. Without communication, it is difficult to prioritize patients plans and make sure the work is done in a timely manner. |  |  | #urgent
77630 | 8/14/2023 4:09:03 PM | Operational/Process Improvement | HSROC, Hazelrig-Salter Radiation Oncology Center | It was discovered while the patient was on the table that there were no shifts on the treat sheet. |  |  | #isocenter
77708 | 8/15/2023 1:53:34 PM | Operational/Process Improvement | TKC-AR, The Kirklin Clinic at Acton Road | Target named PTV_Axilla in the plan does not match the name of the corresponding target in the prescription, 'PTV_Ax.' There was no prior X-check safety timestamp. |  |  | #target
77718 | 8/15/2023 3:01:10 PM | Operational/Process Improvement | UAB Proton Center | Plan not treatment approved or QA in Powerchart at 11:50am for 1:15pm new start. |  |  | #urgent
77736 | 8/15/2023 5:38:08 PM | Operational/Process Improvement | HSROC, Hazelrig-Salter Radiation Oncology Center | Therapist printed out patient's pre-tx chart review paperwork. Patient's yellow sheet was correct, but the Rx included in the paperwork was for another patient. Therapist notified dosimetrist and retrieved the correct prescription from ARIA. |  |  | #documentation
77668 | 8/15/2023 8:14:18 AM | Other Safety Incident | HSROC, Hazelrig-Salter Radiation Oncology Center | Patient did not have gating in RX. The yellow sheet said gating and had a gating setup sheet included. Therapist caught the missing info from the rx and notified physics who had the MD change the rx. |  |  | #rx
77672 | 8/15/2023 8:59:28 AM | Operational/Process Improvement | HSROC, Hazelrig-Salter Radiation Oncology Center | Patient was scheduled for ix. The schedule new start task had ix as the scheduled machine listed by dos in the task. When therapist was doing chart check, it was noticed that the patient was actually planned for the stx. patient had to be rescheduled for stx. |  |  | #machineswitch
77836 | 8/17/2023 10:59:52 AM | Operational/Process Improvement | HSROC, Hazelrig-Salter Radiation Oncology Center | Therapist called pt back for tx and opened chart then realized there was plan revision paperwork in chart. Therapist then looked into RT summary and saw the initial plan was completed early with a revision plan treat approved. Patient had to wait so therapist could check the chart. The ARIA appointment was never changed to Plan Revision (green) so therapist knew a change was coming. |  |  | #revision
77889 | 8/18/2023 9:58:41 AM | Operational/Process Improvement | TKC-AR, The Kirklin Clinic at Acton Road | For an Electron plan, Plan PDF was not uploaded to PowerNote. I was notified last minute that the treatment was set for the next day at 9 am, and the document did not get uploaded til the next morning with less than 1 hour to go. Electron plan checks cannot be completed without referencing the MU calc sheet. On a personal note, I understand the dosimetrist was unable to upload the document until they had access to their work computer, but I'm still submitting this as data for operational improvement. |  |  | #urgent
77071 | 8/2/2023 11:17:28 AM | Operational/Process Improvement | TKC-AR, The Kirklin Clinic at Acton Road | New start task stated to schedule on the True Beam. But it was planned for the IX. |  |  | #machineswitch
78016 | 8/22/2023 10:49:47 AM | Operational/Process Improvement | TKC-AR, The Kirklin Clinic at Acton Road | MIP30-70 was improperly created. The image labeled MIP30-70 shows that 'mean' was selected. When images are blended, it is clear that they are the same image. |  |  | #4dct
78055 | 8/22/2023 9:11:20 PM | Operational/Process Improvement | HSROC, Hazelrig-Salter Radiation Oncology Center | VMAT breast plan was completed and met all constraints for heart. Was not notified that the patient had heart disease or given any constraints prior to planning. Plan was sent back for revision but put in planning approval status. Planner had to duplicate structure set in order to work on revision plan. |  |  | #revision
78109 | 8/23/2023 3:42:32 PM | Near-miss | HSROC, Hazelrig-Salter Radiation Oncology Center | Patient had a cardiac loop recorder implanted near the target, in field (<3 cm from PTV border to structure). The loop recorder was noted in the simulation note and a special physics consult task was added appropriately. However, the dosimetrist missed the loop recorder and did not contour it. The loop recorder was incidentally noticed as the patient was being evaluated for planning on the Ethos. Physics contoured the loop recorder and found that it was receiving a maximum dose of 44 Gy. This was noted around 1:30 pm the day before the patient was scheduled to start (8 am) and a revision for the Eclipse plan was launched to start as soon as possible. There are separate issues with the Ethos-Eclipse workflow which will be logged in a separate event |  |  | #pacemaker
78114 | 8/23/2023 4:16:23 PM | Near-miss | HSROC, Hazelrig-Salter Radiation Oncology Center | Patient did not have any gating instructions in RX. When we pulled the patient up, the gating default pulled up. The patient was planned Breath hold but it was not indicated on the yellow sheet or elsewhere. Patient was verified by physics that it was planned on the breath hold scan. Patient was delayed 10 mins while RX was fixed. |  |  | #Gating
78113 | 8/23/2023 4:23:12 PM | Operational/Process Improvement | HSROC, Hazelrig-Salter Radiation Oncology Center | A patient was originally planned to be treated on the IX and a treatment plan was approved by physician, however due to high lung dose, the MD asked dosimetry/physics to look into creating an Ethos plan to reduce dose via adaption. Due to limited dosimetry coverage, a physicist took over Ethos treatment planning. When the Ethos planning process was initiated, it was noticed that the patient had a cardiac loop recorder that was not contoured/optimized around in the IX plan. Physician was notified; however, the original planning dosimetrist was not notified (physics was operating under the assumption that the patient was only going to be treated on the Ethos). Meanwhile, the normal care path for the IX plan was moved along and the patient was scheduled to start on 8/24. The patient was chemo coordinated so start date was not flexible. The task for physics plan review (on the IX plan) came up the afternoon before the patient was set to start at 8 am. Due to the excessive dose to the loop recorder (see separate incident), physics asked for a plan revision on the IX plan to start as soon as possible |  |  | #machineswitch
78193 | 8/24/2023 5:06:28 PM | Operational/Process Improvement | TKC-AR, The Kirklin Clinic at Acton Road | 800cGy/fx patient. Does not have specification regarding physician supervision in eRx. Therapy happened to notice comment in dosimetry's new start task to get DOD supervision. Directive should be in eRx. |  |  | #documentation
78271 | 8/25/2023 4:40:00 PM | Operational/Process Improvement | HSROC, Hazelrig-Salter Radiation Oncology Center | A patient was simmed using both FB and BH techniques. Target volumes were contoured on both and the autoseg was run on the FB scan. There were no notes in PM on which scan to move forward with. The RX did say BH, but this is not always correct. Physics completed the DOs on the FB scan. Dosimetry asked for clarity on which dataset to use, since they weren't identical (one set of PTVs were different volumes). The MD wanted the BH scan to be used, so the DOs needed to be re-reviewed and adjusted. Potential process improvement could be to have physicians specify in PM what they would like to use prior to planning. Another possible improvement could be deleting out unwanted TVs before turning over volumes for planning. |  |  | #intent
78273 | 8/25/2023 4:50:02 PM | Operational/Process Improvement | HSROC, Hazelrig-Salter Radiation Oncology Center | Patient was simulated on wednesday 8/23/23 but was never moved forward on process maker. Dosimetry did not know that the patient needed to be imported until the physician was asking for it. Is it possible to have simulation inactivate unused "simulation scheduled" status for patients who do not end up getting simulated? It might be easier for Dosimetry to see that is it odd for a patient to have gone past a simulation scheduled date. |  |  | #processmaker
78395 | 8/29/2023 11:36:49 AM | Near-miss | HSROC, Hazelrig-Salter Radiation Oncology Center | TBI sent for reivew with less than 24 hours before tx. It was noted that there was a severe artifact near the elbow, most likely due to motion during the scan and recon. There are multiple failure points: 1) Image Review at CT, on any window/level there is obvious abnormalities on the scan. 2) Dosimetry noticed the issue and asked the physicist for assistance in contouring the body/PTV as well as guidance on what to do with added/missing tissue, the body was fixed but the tissue was not drawn in. This led to a sub-optimal but acceptable dose distribution from the physician since we are on the clock for TBI tx. |  |  | #urgent
78404 | 8/29/2023 12:37:19 PM | Operational/Process Improvement | TKC-AR, The Kirklin Clinic at Acton Road | Treatment field names were not labeled with the correct/actual gantry angles. |  |  | #documentation
78413 | 8/29/2023 2:47:45 PM | Operational/Process Improvement | HSROC, Hazelrig-Salter Radiation Oncology Center | New start SBRT scheduling email sent 08/22. Email requesting update sent 8/29. The person covering was not aware the patient had not been scheduled. |  |  | #scheduling
78433 | 8/29/2023 5:41:35 PM | Unsafe condition | HSROC, Hazelrig-Salter Radiation Oncology Center | While making an SBRT plan, the dosimetrist noticed artifacts that were not accounted for by structures approved by physics during the initial Density Override review. SRS POD was asked to review the structures. After re-contouring and applying revised overrides, dosimetric differences were noticeable. While the gross differences, in this case, were not catastrophic, override errors such as the ones in this case can result in conditions that lead to physicians evaluating plans with inaccurate dosimetry metrics. |  |  | #Density
78446 | 8/30/2023 9:51:46 AM | Near-miss | UAB Proton Center | Physician has drawn 3 target volumes, but only has a prescription for 2 targets. Rx says bilat, volumes are ipsilateral. |  |  | #Rx

## Sheet1
|  |  | Sim order by resident? | Physician | Physician
3/1/2023 10:35:26 AM | TKC-AR, The Kirklin Clinic at Acton Road | 3095404 | 0 | F | Jacob | A
3/2/2023 6:54:39 PM | HSROC, Hazelrig-Salter Radiation Oncology Center | 1231259 | 0 | F | Soike | B
3/24/2023 8:53:46 AM | TKC-AR, The Kirklin Clinic at Acton Road | 2928538 | 0 | G | Boggs | C
4/19/2023 1:12:19 PM | UAB Proton Center | 1145551 |  |  | Spencer | D
5/12/2023 10:17:30 AM | TKC-AR, The Kirklin Clinic at Acton Road | 3959955 |  |  | Keene | E
5/22/2023 7:54:51 AM | HSROC, Hazelrig-Salter Radiation Oncology Center | 3613451 |  |  | Fiveash | F
5/3/2023 10:07:46 AM | UAB Proton Center | 762515 |  |  | Dobie | G
6/20/2023 10:43:11 AM | UAB Proton Center | 4099485 |  |  | Bredel | H
6/28/2023 5:27:32 PM | UAB Proton Center
6/29/2023 5:15:23 PM | UAB Proton Center
6/8/2023 2:01:05 PM | TKC-AR, The Kirklin Clinic at Acton Road | 3927672
6/8/2023 2:01:05 PM | TKC-AR, The Kirklin Clinic at Acton Road | 3927672
6/8/2023 4:06:51 PM | HSROC, Hazelrig-Salter Radiation Oncology Center | 4107108
6/8/2023 4:06:51 PM | HSROC, Hazelrig-Salter Radiation Oncology Center | 4107108
7/28/2023 2:48:30 PM | UAB Proton Center
7/28/2023 4:53:05 PM | TKC-AR, The Kirklin Clinic at Acton Road | 4155130
7/31/2023 12:22:57 PM | TKC-AR, The Kirklin Clinic at Acton Road
7/31/2023 9:19:19 AM | TKC-AR, The Kirklin Clinic at Acton Road | 4122079
8/11/2023 5:45:15 PM | TKC-AR, The Kirklin Clinic at Acton Road | 886093 | 0 | A
8/11/2023 5:41:04 PM | TKC-AR, The Kirklin Clinic at Acton Road | 1797931 | 0 | H